## Slide 1: Nb3Sn Cavity ERL1-4
History of ERL1-4
  Nb3Sn coating, just HPR
    Test shows relatively flat Q vs E up to quench at ~13 MV/m, Q~1e10
  5x HF rinse, HPR
    November test shows Q-slope
  120 K Torture
    Test shows no degradation
  Remove coating with BCP
    Baseline test: good performance up to 20 MV/m
  Recoat with same parameters
    Test shows performance is repeatable

## Slide 3: Measuring DC Flux Penetration Field
Nick made a 1760-winding NbTi solenoid
Get cavity to a few MV/m, then apply DC field, watch Pt (or Pr)
When RF field drops, flux from DC field has penetrated to RF surface
Hall probe measures magnetic field

## Slide 4
Flux penetration
T = 6.3 K

## Slide 5
Approximate Bsh of Nb with correction for field at hall probe
Flux penetration as a function of T2
Lines with x-intercept matching Tc of Nb or Nb3Sn and y-intercepts fitted by eye

## Slide 6
From Nick’s thesis:
Ratio of X’s gives magnetic field correction for hall probe